Project ID: 2017-1598

Lab Hub: Mount Sinai CHEAR Network Laboratory Hub

Project Name: Evaluation of Environmental Exposures in TEDDY

Project PI: Dr. Krischer

Description of Analyses:

Quality control (QC) measures included analysis of the initial calibration, initial calibration verification, and continuing calibration verification standards: NIST traceable mixed-element standard solution at two concentration levels, procedural blanks and repeated analysis of 2% of samples. If QC materials did not meet control requirements, following Westgard multi-rules, the batch was reprocessed and rerun. Matrix-appropriate SRMs and CRMs were analyzed once per study. Results provided were the average of five replicate measurements. Results were corrected for blanks but not for recoveries.

Sample Preparation: samples (200 µL) were diluted 10 ml with diluent solution containing 0.05% Triton X-100, 0.5% Nitric acid and mixed internal standard. Samples were mixed thoroughly and analyzed on the same day.

Instrumentation: inductively coupled plasma – mass spectrometer -Triple quadrupole (ICP-MS) (Agilent 8800-QQQ) and Direct Mercury Analyzer (DMA)-80 Milestone Inc., Monroe, CT)

The daily limit of detection values was calculated as three times the standard deviation of the concentrations found in the method blank for determination of below detection values.

Recoveries were determined from analyses of matrix appropriate CHEAR-QC standards. Recoveries were 90%-110 for As, Ba, Cd, Co, Cr, Cs, Cu, Mo, Ni, Pb, Sb, Sn, Tl, V and Zn; 80%-120% for Mn and 140%-150% for Mg and Al. Data for all QC analysis including Interday CV could be found in Table 2 in the Data Results Reports noted above from DC. Intraday CVs were determined in each batch for internal urine pools fortified at mid- and high-level concentrations analyzed after initial calibration and every ten study samples. Intraday CV range from 0.3 to 21% for all elements.

The laboratory participates in the New York State Department of Health Biomonitoring Proficiency testing program and CTQ The Centre de toxicology du Quebec Multielement External Quality Assessment Scheme.