Analytical Method Summary

Date: 11-19-2019

Project ID: 2017-1762

Project Name: Biological Responses to Tobacco Smoke Exposure in Ill Children: Inflammatory Processes and Oral Metabolomic Profiles

Project PI: Mahabee-Gittens

Summary:

Saliva cotinine was analyzed as previously described with the exception that the analyses were performed on a Luna C18 column1.    Briefly, the methods used were 96-well plate-based liquid chromatography-tandem mass spectrometry assays.   Saliva cotinine was quantified by the area ratio of the analyte to the deuterated standard using specific MS/MS transitions.  Calibration curves were established before each set of LC-MS/MS analyses.  The same sample of deuterated analyte was used for constructing the calibration curve as that added to each sample as internal standard.  Lower limit of quantitation was calculated based on 3 times the noise measured in an extracted blank and was 0.1 ng/mL.

The CHEAR-provided quality control samples had coefficients of variation 2% (pool A) and 3% (pool B).   For Aim 2, the CHEAR-provided quality control samples had coefficients of variation 6% (pool A) and 8% (pool B).

Two pools of diluted smoker’s urine were used as a positive control.  Positive control 1 had a mean of 8.1 ng/mL ±7.8 (3% CV) and positive control 2 had a mean of 50.7 ng/mL ±6.9 (14% CV).

References

(1)	Murphy, S. E., Wickham, K. M., Lindgren, B. R., Spector, L. G., and Joseph, A. (2013) Cotinine and trans 3'-hydroxycotinine in dried blood spots as biomarkers of tobacco exposure and nicotine metabolism. J. Expo. Sci. Environ. Epidemiol 23, 513-518