Project ID: 2017-2121

Lab Hub: The National Exposure Assessment Laboratory at Emory

Project Name: The Dynamics of Exposure, Phthalates and Asthma in a Randomized Trial (DEPART)

Project PI: Dr. Karr

Description of Analyses:

Urine samples were randomized prior to analysis to reduce the amount of analytic bias introduced into the data.  A 1-mL aliquot of urine was spiked with isotopically labeled analogues of the target pesticide metabolites then was subjected to an enzyme hydrolysis to liberate glucuronide-bound conjugates.  The hydrolysate was extracted using an OASIS solid phase extraction column, washing with a an acidic 5% methanol solution and eluting with methanol.  The extract was concentrated to dryness and reconstituted in mobile phase for analysis using liquid chromatography-tandem mass spectrometry (LC-MS/MS).  Analyte concentrations were calculated using isotope dilution calibration.  Two quality control materials (one high and one low) and one blank sample were analyzed concurrently with each set of 28 unknown samples.  Further quality assurance measures were included in the sample analyses including the analysis of NIST SRM 3672 and 3673 (one of each per 100 samples), inclusion of blinded quality controls (6 each of 2 materials per 100 samples), and bi-annual participation in the German External Quality Assessment Scheme (G-EQUAS).  Specifications of the method are provided in Table 1.

Table 1.  Target analytes and method specifications (ng/mL)

| Acquisition Method | Target Analyte | Limit of Detection* | Total RSD (%)* | Intraday RSD (%)* | Interday RSD (%)* | Accuracy** |
| --- | --- | --- | --- | --- | --- | --- |
| UPM | 3,5,6-trichloro-2-pyridinol | 0.5 | 10 | 4 | 10 | 101 |
|  | 3-phenoxybenzoic acid | 0.5 | 10 | 7 | 10 | 98 |

*Values will vary across studies and samples. RSD=relative standard deviation

** Accuracies are reported when certified reference materials are available for analysis.  Otherwise, relative recoveries (also called spiked recoveries) are reported.  They are calculated as the (quantified value/expected value)*100.