Project ID: 2017-2121

Lab Hub: Mount Sinai CHEAR Network Laboratory Hub

Project Name: The Dynamics of Exposure, Phthalates and Asthma in a Randomized Trial (DEPART)

Project PI: Dr. Karr

Description of Analyses:

We followed the analytical method of the Centers for Disease Control and Prevention for phthalate metabolites in urine (CDC) (Ye et al., 2005) with minor modifications (Gavin et al., 2014). Quantification was based on an isotope-dilution liquid chromatography and tandem mass spectrometry method. In brief, 13C12, 13C6 or D6 labeled internal standards were added to each sample; metabolites were deconjugated with β-glucuronidase /arylsulfatase from Helix pomatia (product # 10127060001, Roche Diagnostics through Sigma Aldrich), followed by solid-phase extraction with an Oasis HLB hydrophilic-lipophilic balanced reversed-phase 96-well plate (30 mg sorbent per well, 30 µm particle size; Waters Corporation, Milford, MA). The procedure was automated using a liquid handler (epMotion 5075vtc; Eppendorf, Hauppauge, NY). The LC-MS/MS (UHPLC Nexera XR, Shimadzu and Sciex 6500 triple quadruple MS, AB Sciex; Framingham, MA) was operated in electrospray negative mode for ionization and multiple reaction monitoring (MRM) for quantification. Chromatographic separation was achieved on a Thermo Betasil C18, 5.0 µm, 100 x 2.1 mm analytical column with 10 x 2.1 mm guard cartridge (Thermo scientific, Waltham, MA) using a mobile phase gradient with LC-MS grade water and acetonitrile, respectively. We analyzed benzophenone-3 (BP3), bisphenol A (BPA), bisphenol F (BPF), bisphenol S (BPS), triclosan (TCS), triclocarban (TCC), methyl Paraben (MEPB), butyl paraben (BUPB), ethyl paraben (ETPB), and propyl paraben (PRPB). We expanded the phenols method to include the hydroxylated urinary metabolites of polycyclic aromatic hydrocarbons (PAHs) [1-hydroxynaphthalene (NAP1), 2-hydroxynaphthalene (NAP 2), and 1-hydroxypyrene (PYR1)] (Schmidt et al., 2013). The limits of detection for environmental phenols and hydroxyl-PAH metabolites in water blanks ranged from 0.05 to 1.0 ng/mL.

Quality controls (QC) included in each batch were procedural and instrumental blanks, matrix spikes in the lower, middle and upper range of assay validation, enzyme de-conjugation and no-enzyme matrix treatments, NIST standard reference material (SRM 3672: Organic Contaminants in Smokers' Urine and SRM 3673: Organic Contaminants in Non-Smokers' Urine), and archived proficiency testing material. Batchwise relative standard deviations (RSDs) of QCs during analysis of the study specimens were <10% for target analytes in reference or fortified material, except for analytes at or below LOQ (LOQ = 3xLOD) where RSDs up to 30% were accepted. Intra-batch precision (CV) was below 10% and inter-batch precision was below 20% for QC analytes above the LOQ. Recoveries in batch QC were between 80% and 120%. We participated in proficiency testing programs for environmental phenols, hydroxylated PAHs, and creatinine conducted by G-EQUAS (http://www.g-equas.de/), OSEQAS (https://www.inspq.qc.ca/en/ctq/eqas/oqesas/), and CHEAR inter-laboratory Round Robin.

References:

- Ye X, Kuklenyik Z, Needham LL, Calafat AM. Automated on-line column-switching HPLC-MS/MS method with peak focusing for the determination of nine environmental phenols in urine. Anal Chem. 2005. 77(16):5407-13.

- Gavin QW, Ramage RT, Waldman JM, She J. Development of HPLC-MS/MS method for the simultaneous determination of environmental phenols in human urine, International Journal of Environmental Analytical Chemistry. 2014. 94 (2): 168-182.

- Schmidt L, Müller J, Göen T. Simultaneous monitoring of seven phenolic metabolites of endocrine disrupting compounds (EDC) in human urine using gas chromatography with tandem mass spectrometry. Anal Bioanal Chem. 2013. 405(6):2019-29.